| Accuracy |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  | Model 14 | Model 15 | Model 16 | Model 20 | Model 21 |
| DC Ranges | ±2% of fsd | ±1.5% of fsd | ±1% of fsd | ±1% of fsd | ±1% of fsd |
| AC Ranges | ±2.5% of fsd | ±2.25% of fsd | ±1.5% of fsd | ±1.5% of fsd | ±1.5% of fsd |
| Sensitivity | DC ranges 20,000 ohms per volt AC ranges 2,000 ohms per volt |
| Ranges |  |  |  |  |  |
|  | Model 14 | Model 15 | Model 16 | Model 20 | Model 21 |
| Direct Voltage | 8 ranges | 8 ranges | 8 ranges | 9 ranges | 9 ranges |
|  | 100mV | 100mV | 100mV | 120mV | 120mV |
|  | 1.0V | 1.0V | 1.0V | 1.2V | 1.2V |
|  | 3.0V | 3.0V | 3.0V | 3.0V | 3.0V |
|  | 10V | 10V | 10V | 12V | 12V |
|  | 30V | 30V | 30V | 30V | 30V |
|  | 100V | 100V | 100V | 120V | 120V |
|  | 300V | 300V | 300V | 300V | 300V |
|  | 1000V | 1000V | 1000V | 600V | 600V |
|  |  |  |  | 1200V | 1200V |
| Direct Current | 10 ranges | 10 ranges | 10 ranges | 9 ranges | 9 ranges |
|  | 0.1mA | 0.1mA | 0.1mA | 120μA | 120μA |
|  | 0.3mA | 0.3mA | 0.3mA | 600μA | 600μA |
|  | 1.0mA | 1.0mA | 1.0mA | 1.2mA | 1.2mA |
|  | 3.0mA | 3.0mA | 3.0mA | 3.0mA | 3.0mA |
|  | 10mA | 10mA | 10mA | 12mA | 12mA |
|  | 30mA | 30mA | 30mA | 60mA | 60mA |
|  | 100mA | 100mA | 100mA | 300mA | 300mA |
|  | 300mA | 300mA | 300mA | 1.2A | 1.2A |
|  | 1.0A | 1.0A | 1.0A | 3.0A | 6.0A |
|  | 3.0A | 3.0A | 3.0A |  |  |

|  | Model 14 | Model 15 | Model 16 | Model 20 | Model 21 |
| --- | --- | --- | --- | --- | --- |
| Resistance | 4 ranges | 4 ranges | 4 ranges | 4 ranges | 4 ranges |
|  | Ω 10 0 - 10kΩ 120Ω centre scale | Ω 10 0 - 10kΩ 120Ω centre scale | Ω 10 0 - 10kΩ 120Ω centre scale | Ω 10 0 - 10kΩ 100Ω centre scale | Ω 10 0 - 10kΩ 100Ω centre scale |
|  | Ω 0 - 100kΩ 1.2kΩ centre scale | Ω 0 - 100kΩ 1.2kΩ centre scale | Ω 0 - 100kΩ 1.2kΩ centre scale | Ω 0 - 100kΩ 1.0kΩ centre scale | Ω 0 - 100kΩ 1.0kΩ centre scale |
|  | Ω 10 0 - 1MΩ 12kΩ centre scale | Ω 10 0 - 1MΩ 12kΩ centre scale | Ω 10 0 - 1MΩ 12kΩ centre scale | Ω 10 0 - 1MΩ 10kΩ centre scale | Ω 10 0 - 1MΩ 10kΩ centre scale |
|  | Ω 1000 0 - 100MΩ 1.2MΩ centre scale | Ω 1000 0 - 100MΩ 1.2MΩ centre scale | Ω 1000 0 - 100MΩ 1.2MΩ centre scale | Ω 1000 0 - 100MΩ 1.0MΩ centre scale | Ω 1000 0 - 10MΩ 100kΩ centre scale |
| Alternating Voltage | 6 ranges | 7 ranges | 7 ranges | 8 ranges | 8 ranges |
|  |  | 1.0V | 1.0V | 1.2V | 1.2V |
|  | 3.0V | 3.0V | 3.0V | 3.0V | 3.0V |
|  | 10V | 10V | 10V | 12V | 12V |
|  | 30V | 30V | 30V | 30V | 30V |
|  | 100V | 100V | 100V | 120V | 120V |
|  | 300V | 300V | 300V | 300V | 300V |
|  | 1000V | 1000V | 1000V | 600V | 600V |
|  |  |  |  | 1200V | 1200V |
| Alternating Current | None | 6 ranges | 6 ranges | 6 ranges | 6 ranges |
|  |  | 3mA | 3mA | 3mA | 3mA |
|  |  | 10mA | 10mA | 12mA | 12mA |
|  |  | 30mA | 30mA | 30mA | 30mA |
|  |  | 0.1A | 0.1A | 0.1A | 0.1A |
|  |  | 0.3A | 0.3A | 1.2A | 1.2A |
|  |  | 1A | 1A | 3A | 6A |
|  |  | 3A | 3A |  |  |
| No. of ranges | 28 ranges | 36 ranges | 36 ranges | 36 ranges | 36 ranges |

Control - all models;

Main Switch - Range value selection and OFF.

Reverse Moving Coil Switch - Button at the centre of the main switch

DC & Ohms or AC - pair of "radio" (alternate) push buttons

Cutout reset - Push to reset, button at left of front panel

Ohms zero adjust - knurled knob at right of front panel

Movement Zero Adjuster - screwdriver adjustment accessed through hole at bottom centre of front bezel

Scales

Model 14

Model 15	Current & Voltage - 0 to 100, figures at intervals of 10, 20, etc.

0 to 30 figures at intervals of 5, 10, 15, etc.

Resistance 		- 0 to 100kΩ, non-linear, zero at rhs.

Anti-parallax mirror

Model 16

Model 20